Summary:

We received 2 LT-GaAs antennas from Hong Zhi. They were labeled 20 1-2 and 40 3-4 with gap sizes 20 um and 50 um (maybe 40 um?) respectively. We tested only 40 3-4 as an emitter since it was most appropriate.

Comparing the performance of our Lithuania LT-GaAs emitters to the 40 3-4 and our PCA-LTGaAs-E, the 40 3-4 has approximately 1/4 the amplitude and our PCA-LTGaAs-E has approximately 1/3 the amplitude.

When making the comparison, only the emitter and emitter lens was optimized to ensure that there was no change in alignment. The 40 3-4 and PCA-LTGaAs-E emitters were tested different days.

Experiment Parameters:

| Setup / System | Test Bed System |
| --- | --- |
| Laser | IMRA CX-20; 23 mW; <100 fs; 7.5 mW at emitter; 4.6 mW after ZnTe |
| THz Emitter | Lithuania 50 um; 40 3-4 (LT from Hong Zhi); PCA-LTGaAs-E |
| Bias [V] | 75 V |
| THz Receiver | 1mm ZnTe; no lens |
| Delay Stage | Nanomotion Stage; Allmotion HR4 Controller; 0.1 um Linear Encoder; 0.5” retroreflector (2x fold) |
| Acquisition | Lock-in |
| Modulation Frequency | 100 kHz |
| Time Constant | 30 ms |
| Measurement Type | Collimated Transmission |
| Emitter OAP FL | 1” |
| Detector OAP FL | 3” |

Figure 1: (a) Plot of the THz time-domain waveforms from the 40 3-4 and Lithuania emitter. Waveforms are normalized (b) to compare pulse shapes. The 40 3-4 emitter is a factor of ~4.0 lower in amplitude compared to the Lithuania emitter.

Figure 2: Plot of the THz frequency-domain spectra from the 40 3-4 and Lithuania emitter. The 40 3-4 is about 12 dB lower than the Lithuania across the entire spectrum (except the low frequencies) which leads us to believe that alignment error is negligible.

Figure 3: (a) Plot of the THz time-domain waveforms from the PCA-LTGaAs-E and Lithuania emitter. Waveforms are normalized (b) to compare pulse shapes. The PCA-LTGaAs-E emitter is a factor of ~3.4 lower in amplitude compared to the Lithuania emitter.

Figure 4: Plot of the THz frequency-domain spectra from the PCA-LTGaAs-E and Lithuania emitter.

# Text boxes

(b)

(a)

(b)

(a)